## Slide 1: Master Silica Sampling Accuracy
Understanding the Dynamics of Size Selective Samplers and Why the PPI Makes Sense
Dusty Ott, CIH

## Slide 2: Learning Objectives:
Understand how size-selective samplers work and why they matter for silica sampling
Learn what makes the SKC PPI unique and how it performs compared to traditional cyclones
Explore common mistakes in silica sampling—and how to avoid them
Break down the true cost of sampling, including the risk of “bad data”
Gain practical tips for successful field use and how samples are analyzed

## Slide 3: Outline
Why Silica Accuracy Matters
Understanding Size Selective Sampling
The SKC PPI: What Sets it Apart
Do’s and Don’ts of Silica Sampling
Affects of Using the Wrong Flow Rate
The True Cost of Sampling

## Slide 4: The Invisible Danger
Lets zoom in…

## Slide 5: Delayed harm from Silica
Respirable crystalline silica is microscopic in size and has a sharp, crystalline structure
These small particles travel deep into the lungs, causing inflammation and scarring of lung tissue.
Prolonged exposure can result in debilitating lung diseases such as silicosis and chronic obstructive pulmonary disease (COPD). This can take years or decades of exposure.

## Slide 6: What is Size Selective Sampling
Size-selective sampling is the process of collecting airborne particles in a way that selectively captures only those within a specific aerodynamic size range, based on the health effect of concern.

## Slide 7: Respirable Crystalline Silica (RCS)
Cannot be seen by the naked eye
RESPIRABLE CRYSTALLINE SILICA
1-10 um
Will travel deep into the lungs

## Slide 8: Collection Efficiency and Cut Point
Collection efficiency refers to the percentage of airborne particles of a particular aerodynamic diameter that are captured for analysis
Smaller particles are collected at a higher efficiency and larger particles are collected at a lower efficiency
The ISO 7708:1995 Respirable curve has a collection efficiency of 50% at 4um
  This is referred to as the “Cut Point”
Each cyclone or impactor has a specific flow rate that must be used to achieve a 50% cut point at 4 um

## Slide 9: Collection Efficiency and Our Bodies
The ISO 7708:1995 curve closely follows what our body naturally does to remove larger particles from entering our lungs
Our body has several natural impaction or deposition sites:
  Nose/Mouth
  Back of throat
  Bottom of Trachea
  Bronchi and Bronchioles
We are our own ‘Oiled Impactor’

## Slide 10: Size selective samplers

## Slide 11: Basic Principles of Cyclone Operation
Inlet Flow: Contaminated air enters the cyclone separator at high speed through a tangential inlet.
Vortex Formation: The air spirals downward in a circular motion inside the cyclone, creating a strong centrifugal force.
Particle Separation: The centrifugal force causes heavier particles to move outward toward the cyclone walls. These particles lose momentum, slide down the walls, and are collected in a hopper at the bottom.
Clean(er) Air Exit: The cleaned air, now free of heavier particulates, moves inward to the center of the cyclone and exits through an outlet pipe called a vortex finder at the top.

## Slide 12: Factors Influencing Cyclone Collection Efficiency
Particle Size Distribution
  Cyclones are not very efficient at collecting small particles (this works in our favor)
Design Parameters
  Diameter, length, etc.
Inlet Flow Rate and Velocity
  Too high can cause turbulence and/or re-entrainment of particles
  Too low will have insufficient centrifugal force to separate particles effectively
Operating Conditions
  Temperature and Humidly
  High/low concentration of dust in the air
Maintenance and Wear
  Cyclone walls can be loaded with dust if not cleaned, changing the internal geometry and flow
  Blockages of the inlet affecting flow

## Slide 13: What’s Inside?

## Slide 14: Basic Principles of Impactor Operation
Air Intake: Air is drawn into the impactor through an inlet.
Velocity Increase: As the air passes through one or more nozzles, it accelerates, increasing the air velocity.
Impaction Plate: The accelerated air stream directs particles towards an impaction plate positioned in the path of the flow. Larger particles, due to their greater inertia, cannot follow the air stream and impact on the plate, where they are collected.
Smaller Particles: Smaller particles, with less inertia, remain in the air stream and are carried to subsequent stages of the impactor or exit the device, depending on the design.

## Slide 15: Principle of Operation
  Instead of a cyclone effect, the PPI uses impactors to remove larger particles
  Contains 4 internal, pre-oiled plates (oiled substrates)
  The impactor plate assembly is sonically welded into place by SKC and requires no assembly
Notes: What is a parallel particle impactor or PPIs?
PPIs are impaction-based filter samplers that perform precise size selection for either thoracic or respirable dust, depending on the model.
-PPI Samplers contain four small impactors in the inlet section of the device.
-Each impactor features a unique 50% cut-point to target a specific one-quarter segment of the ISO/CEN curve resulting in a precise fit along the entire curve.
-A sample pump operating at 2, 4, or 8 L/min (2 L/min only for thoracic) pulls air through the inlet nozzle of each impactor in the inlet plate.
-Particles larger than each impactor’s 50% cut-point are scrubbed and retained by impaction onto the porous oiled impaction substrate contained in each impactor.
-Smaller particles continue to the standard 37-mm collection filter for analysis.

## Slide 16: Disposable PPI
8 L/min
4 L/min
2 L/min
Full shift sampling
Task-based sampling or low dust environments
Notes: Discuss the slides.
The thoracic PPI sampler is used for sampling metalworking fluids using the NIOSH method 5524, hard metals with Cobalt and tungsten carbide as cobalt.
Anti-static plastic

## Slide 17: Advantages Over Traditional Cyclone
A calibration adapter is available to attach the PPI to the flowmeter (previously know as “Calibrator”).
No tipping hazard. You can invert the sampler without causing large particles to land on the filter invalidating the sample.
No cleaning necessary after the sampling
Much smaller and easier to wear
Workers prefer wearing the PPI compared to a cyclone
Readily available through RAECO Rents

## Slide 18: Full Regulatory Compliance
Approved for use in both the OSHA and MSHA Final Silica rule

## Slide 19: What do I need to Sample for Silica?
Constant flow pump
Flow verification device
Size selective sampler
  Tygon tubing
  Cassette holder if using a Cyclone
Pre-weighed 5um PVC filter
  Paired with 37mm Cassette or PPI
Note Pad and pens
Camera/Phone if allowed
Equipment is available to rent from RAECO

## Slide 21: Proper Placement

## Slide 22: Don’ts of Silica Sampling
In General:
  Not doing a pre/post flow rate verification: Without both checks, you can’t confirm sampling accuracy, and results may be invalid.
  Not taking notes in the field: Missing task descriptions, environmental conditions, and anomalies makes interpreting results difficult.
  Not sampling in the breathing zone: Samples not captured in the breathing zone do not accurately depict personal exposure
  Not using an accredited lab for analysis: Labs without proper accreditation may not follow recognized quality standards, putting your data reliability at risk.

## Slide 23: Don’ts of Silica Sampling
Cyclones:
  Inverting the cyclone: Large particles in the grit pot have the ability to migrate to the collection media
  Cyclone inlet not facing the correct way: Inlet needs to be properly oriented, or it may be blocked by user
  Not cleaning after each use: Dust can collect and accumulate on the walls of the cyclone effectively reduce the diameter of the cyclone. This can lead to a reduced cut-point
  Incorrect flow rate: Each cyclone has a specific flow rate. Increasing or decreasing the flow directly impacts the size of particles collected
  Operating without a grit pot: Air will enter both from the top and bottom of the cyclone disrupting the airflow and voiding the sample
  Not using multiple cassettes when sampling high dust tasks: This can lead to filter overloading and additional costs for analysis. No more than 2mg of dust on the filter per NIOSH 7500

## Slide 24: Don’ts of Silica Sampling
PPI:
  Incorrect flow rate: Each PPI has a specific flow rate. Increasing or decreasing the flow directly impacts the size of particles collected
  Not using multiple PPI’s when sampling high dust tasks: Oiled impactors can handle between 6-7mg of dust before becoming overloaded and ‘spilling’ onto the filter.
    This can lead to filter overloading and additional costs for analysis. No more than 2mg of dust on the filter per NIOSH 7500

## Slide 25: Flow Rates and Their Impact on Cut Point
Flow rate in inversely related to the cut point
  Higher flow rate results in a lower cut point
  Lower flow rate results in a higher cut point

## Slide 26: SKC Aluminum Cyclone
SKC Aluminum Cyclone

## Slide 27: Practical Representation

## Slide 28: What Happens After a Sample is Taken?
Post - flow rate verification is performed
Start/Stop time are noted and used to calculate volume
  (total minutes x flow rate = volume)
  600 minutes x 2 L/min = 1200 L (1.2m³)
Lab Chain of Custody (COC) is filled out and shipped with samples to the lab
ACME Analysis

## Slide 29: How are Samples Analyzed?
Samples are received by the lab
ACME Analysis

## Slide 30: How are Samples Analyzed?
Pre-weighed filter is removed from PPI or 37mm cassette and weighed again
  This is why using a pre-weighed filter is a necessity
  Analytical error between PPI and 37mm cassette is the same
    Both are analyzed exactly the same

## Slide 31: How are Samples Analyzed?
After weighing, filter is then digested in tetrahydrofuran

## Slide 32: How are Samples Analyzed?
Tetrahydrofuran/dust solution is then filtered through a silver membrane filter with a funnel and vacuum pump

## Slide 33: How are Samples Analyzed?
Silver membrane filter is then placed into the XRD Analyzer to determine silica content
  Each polymorph has a different ‘signature’

## Slide 34: How Much Does Analysis Cost?
NIOSH 0600 by Microbalance
  Respirable Dust - $24
NIOSH 7500 Silica by XRD*
  Quartz - $75
  Quartz + Cristobalite - $85
  Quartz + Cristobalite + Tritymite - $95
Media Costs
  ~$5 Pre-weighed 37mm Cassette
    $90 for each Cyclone
  ~$29 Pre-weighed PPI
*Samples having particulate weights that exceed the method-specified maximum filter loading will be subject to surcharges for additional sample preparation and analysis fees. Filters with particulate weights exceeding 2mg will be assessed an additional $25 (up to 2x additional sample splits at $25 each).
Filters with particulate weights exceeding 7mg will be processed as a billable bulk sample at the applicable bulk prices.
  Quartz - $160
  Quartz + Cristobalite - $175
  Quartz + Cristobalite + Tritymite - $190
Prices provided by ALS in SLC, UT

## Slide 35: Additional Costs to Consider
Resampling due to invalid samples:
  Consider the additional time required for you, the supervisor, as well as the person being sampled
Follow-up sampling based on initial results
  Frequency will be determined by initial results
Implementing Controls
  Elimination
  Substitution
  Engineering
  Administrative
  PPE

## Slide 36: Additional Costs to Consider
Regulatory penalties for overexposures
Labor costs per sample:
  MSHA estimates that M/NM mines with 21-500 employees have a labor cost of $380 per sample
  This does not include the cost of production loss from supervisors and miners during sampling

## Slide 37: Reasons to Use a PPI
Labor time saved
  Faster setup/teardown, no cleaning
Fewer voids/re-samples
  Less re-collection, re-shipping, and re-reporting
Higher worker acceptance
  Preferred by employees, smaller, and less intrusive
No equipment maintenance
  No cleaning, no repairs or spare parts needed
Convenience
  Overall easier to prepare, deploy, and sample with
  Makes your life easier
  Actual picture of an IH after a full-shift sample

## Slide 38: THANK YOU!
Here’s how to contact me for more information!
E: dott@skcinc.com
P: 724.678.1096
www.skcinc.com
FOLLOW SKC!
https://www.raecorents.com/blog/
www.linkedin.com/company/raeco-rents
www.youtube.com/user/RaecoRentsVideos
Blog
Catch up on previous webinars at:
https://www.raecorents.com/training-center/webinars